ECG SOFTWARE

TECHNICAL AUDIT REPORT

Modular 12-Lead ECG Analysis System

| Review Date | February 12, 2026 |
| --- | --- |
| System Language | Python 3.x |
| Sampling Rate | 500 Hz (12-bit ADC) |
| Leads Supported | 12-Lead (I, II, III, aVR, aVL, aVF, V1-V6) |
| Reviewer | Claude (Anthropic) — AI Technical Audit |

# 1. Executive Summary

This report presents a line-by-line cross-verification of the Modular 12-Lead ECG Analysis System against industry-standard clinical ECG processing practices. Every major parameter — HR, PR, QRS, QT/QTc, P wave — has been reviewed for algorithmic correctness, numerical stability, and real-world monitoring reliability.

Overall, the system demonstrates a solid engineering foundation with a correct filter chain, a thoughtful multi-strategy R-peak detection approach, and important fixes already applied (HR smoothing, asystole guard, variance normalization removal). However, several significant gaps remain that affect clinical reliability in real-world monitoring conditions.

Overall Score by Domain

| Domain | Score | Summary |
| --- | --- | --- |
| Heart Rate (HR) | 78 / 100 | Good multi-strategy detection. EMA smoothing applied. Missing ectopic beat rejection. |
| Signal Filtering | 85 / 100 | Correct filter chain order. Gold-standard baseline wander removal. Filter order (DFT→EMG→AC) should be reordered. |
| PR / QRS Intervals | 58 / 100 | PR uses P-to-Q method (correct) but PR filter range too narrow. QRS uses Q-to-S (correct) but Q search window too short. Critical bugs present. |
| QT / QTc | 42 / 100 | Most critical section. T-peak used instead of T-end (fundamental error). No T-end detection = inaccurate QT. QTc formula is correct but input is wrong. |
| P Wave Detection | 55 / 100 | Search window (120-200 ms before R) is correct. P duration calculation has a code bug (peak_rel variable undefined). 20% amplitude threshold is too high for low-amplitude P waves. |
| Lead-Off Detection | 45 / 100 | Recommended implementation provided but NOT integrated into live calculation pipeline. If a lead is off, metrics are still calculated from noise. High-variance threshold (10,000) is too aggressive. |
| Value Stability | 60 / 100 | HR is now stabilized. PR, QRS, QT, P duration have NO smoothing or beat-rejection gating — will flicker on any real patient signal. |

# 2. Critical Bugs Found — Fix Immediately

🔴 BUG #1 — QT Interval: T-Peak Used Instead of T-End This is the most clinically dangerous error in the codebase. QT is measured from QRS onset to the END of the T wave. Your code uses T-peak as the endpoint. T-peak typically occurs ~100-150 ms before T-end. This means your QTc will be systematically ~80-120 ms too short, making dangerously prolonged QTc intervals appear normal.

Current Code (intervals.py — WRONG):

# T-peak: 200-400 ms after R

t_start = r_idx + int(0.2 * sampling_rate)

t_end   = r_idx + int(0.4 * sampling_rate)

t_segment = filtered_signal[t_start:t_end]

t_idx = t_start + np.argmax(t_segment)   # <-- THIS IS T-PEAK, NOT T-END

# QT = t_idx - qrs_onset   <-- WRONG, understates QT by ~100ms

Fixed Code — Tangent Method for T-End:

def find_t_end_tangent(signal, t_peak_idx, sampling_rate=500):

"""Find T-wave end using the tangent-intersection method."""

# Search from T-peak up to 600ms after R (t-peak offset ~150-200ms)

search_end = min(len(signal)-1, t_peak_idx + int(0.25 * sampling_rate))

segment = signal[t_peak_idx:search_end]

if len(segment) < 10:

return t_peak_idx  # Fallback

# Find steepest descending slope after T peak

deriv = np.gradient(segment)

steepest = np.argmin(deriv)  # Most negative slope

# Draw tangent line from steepest slope to baseline

slope = deriv[steepest]

baseline = np.median(signal[t_peak_idx + int(0.3*sampling_rate):

t_peak_idx + int(0.4*sampling_rate)])  # Isoelectric

# x-intercept = peak_val - baseline / slope

peak_val = segment[steepest]

if abs(slope) < 1e-6:

return t_peak_idx + steepest

dx = (peak_val - baseline) / abs(slope)

t_end = t_peak_idx + steepest + int(dx)

t_end = min(t_end, search_end)

return max(t_peak_idx, t_end)

# Then QT calculation becomes:

t_peak_idx = t_start + np.argmax(t_segment)

t_end_idx  = find_t_end_tangent(filtered_signal, t_peak_idx, sampling_rate)

qt_ms = (t_end_idx - qrs_onset_idx) / sampling_rate * 1000

🔴 BUG #2 — P Duration: peak_rel Variable is Undefined In calculate_p_duration() (expanded_lead_view.py:2259-2295), the variable peak_rel is used in the onset/offset search loop but is never defined. This will throw a NameError at runtime whenever P duration is calculated. The code will silently fall through to the except block, returning 0 for P duration every time.

Current Code (BROKEN):

# Find onset and offset

left = peak_rel     # <-- NameError: peak_rel is not defined!

while left > 0 and np.abs(segment[left] - baseline) > thresh:

left -= 1

Fixed Code:

# Correct: compute peak_rel from segment array

peak_rel = np.argmax(np.abs(segment - baseline))  # position within segment

left = peak_rel

while left > 0 and np.abs(segment[left] - baseline) > thresh:

left -= 1

right = peak_rel

while right < len(segment)-1 and np.abs(segment[right] - baseline) > thresh:

right += 1

dur_samples = max(1, right - left)

p_durations.append(dur_samples * 1000.0 / sampling_rate)

🟠 BUG #3 — QRS Filter Range Excludes Normal Values In calculate_qrs_duration(), the physiological filter is set to 60-120 ms. But the AHA standard for normal QRS is 70-110 ms, and many borderline BBBs fall at 110-120 ms. More importantly: for pediatric patients QRS can be 50-60 ms. Using 60 ms as a hard floor will silently return 0 for narrow pediatric QRS. Consider widening to 40-160 ms and letting the clinical interpretation layer handle flagging.

🟠 BUG #4 — PR Interval Filter Too Narrow (Rejects Normal Short PR) The PR filter range is hard-coded to 120-200 ms. This silently discards any PR interval below 120 ms, meaning pre-excitation (WPW syndrome, delta waves with PR ~80-110 ms) will return PR = 0 instead of a short value. Change the filter floor from 120 to 80 ms and flag short PR as a finding rather than discarding it.

🟠 BUG #5 — Q-Wave Search Window Too Short (40 ms) Q-wave search is set to 40 ms before R-peak. In patients with bundle branch block, inferior infarction, or pathological Q waves, the Q wave onset can begin 60-80 ms before R. Widen to 60 ms (int(0.06 * fs)) to capture these cases, otherwise QRS duration will be underestimated.

# 3. Parameter-by-Parameter Deep Audit

## 3.1 Heart Rate (HR)

The HR calculation is the most mature part of the codebase and the approach is largely correct. The multi-strategy peak detection is a good design — running three strategies and selecting the one with lowest RR interval standard deviation is sound engineering. The EMA + median buffer smoothing is correctly implemented.

What is correct:

- Multi-strategy detection (conservative/normal/tight) with automatic selection by lowest std deviation is industry-aligned.

- Physiological range filter (200-6000 ms RR = 10-300 BPM) is correct.

- Median RR interval for HR calculation is the right approach (vs mean, which is distorted by ectopic beats).

- EMA alpha = 0.5 for large changes and 0.10 for small changes is a good adaptive approach.

- Buffer size of 15 samples with median calculation correctly rejects outlier beats.

What is missing:

- No beat type classification before RR averaging. PVCs (premature ventricular contractions) produce very short RR followed by a compensatory long RR. Without excluding V-beats, HR will be inaccurate during PVCs.

- The 0.5 * std height threshold will fail on inverted leads (e.g., aVR) where R is negative. The code operates only on Lead II which is correct, but if Lead II is accidentally negative-dominant, all HR will fail.

- No check for T-wave double-counting. At high HR (>150 BPM), tall T waves can be detected as R peaks. Adding a check that successive peak amplitudes match within 2x would catch this.

Recommended additional fix:

# After valid_intervals is computed, add ectopic rejection:

median_rr = np.median(valid_intervals)

# Reject any RR interval deviating more than 20% from median (ectopic beat filter)

normal_intervals = valid_intervals[np.abs(valid_intervals - median_rr) < 0.20 * median_rr]

if len(normal_intervals) >= 2:

valid_intervals = normal_intervals

heart_rate = 60000 / np.median(valid_intervals)

## 3.2 PR Interval

The PR calculation method — matching P peaks to Q peaks and computing the time difference — is the standard approach. However, two issues affect accuracy significantly.

Issue 1 — P onset vs P peak:

Clinical PR interval is measured from the onset (beginning) of the P wave to the onset of the QRS complex, not from P peak. Using P peak instead of P onset systematically overstates PR by approximately 40-60 ms (half the P wave duration). Similarly, using Q peak instead of QRS onset understates the PR endpoint. Both errors partially cancel but introduce variable error depending on P wave morphology.

Recommended approach:

# Find P wave onset (left threshold crossing, already computed in P duration)

# Find QRS onset (first departure from baseline before Q peak)

def find_qrs_onset(signal, q_idx, sampling_rate=500):

search_back = int(0.05 * sampling_rate)  # 50ms before Q

start = max(0, q_idx - search_back)

segment = signal[start:q_idx]

baseline = np.median(signal[start-int(0.1*sampling_rate):start])

thresh = 0.1 * np.abs(signal[q_idx] - baseline)

for i in range(len(segment)-1, 0, -1):

if np.abs(segment[i] - baseline) < thresh:

return start + i

return start

Issue 2 — P wave search window:

The P wave search window of 120-200 ms before R-peak is reasonable for sinus rhythm at normal HR. However at high HR (>100 BPM, RR < 600 ms), the P wave can occur only 80-100 ms before R. Extend the search window to 80-250 ms before R to capture P waves at elevated heart rates.

## 3.3 QRS Duration

The Q-to-S method for QRS duration is fundamentally correct. S-wave detection (minimum in 80 ms window after R) is well-chosen. The main issues are the Q-search window (too short, as noted in Bug #5) and the hard physiological filter range.

Additional concern — Baseline return vs S peak:

S-peak detection finds the minimum amplitude in 80 ms after R. This is the S-wave amplitude minimum, not the QRS end (J-point). In patients with wide QRS (bundle branch block), the S wave can be wide and the nadir does not coincide with the return to baseline. For clinical-grade QRS width, the endpoint should be where the signal returns to isoelectric baseline (J-point), not just the S nadir. This is a secondary issue but important for BBB detection accuracy.

# Better S-end / J-point detection:

def find_j_point(signal, r_idx, sampling_rate=500):

search_end = r_idx + int(0.2 * sampling_rate)  # Up to 200ms after R

segment = signal[r_idx:search_end]

baseline = np.median(signal[r_idx + int(0.25*sampling_rate):

r_idx + int(0.35*sampling_rate)])

thresh = 0.1 * np.std(segment)

for i in range(len(segment)):

if abs(segment[i] - baseline) < thresh:

return r_idx + i

return r_idx + int(0.1 * sampling_rate)  # Fallback

## 3.4 QT / QTc Interval — Most Critical Section

The QTc formula (Bazett) is mathematically correct. The clamping to 200-600 ms is appropriate. However the fundamental input is wrong (see Bug #1 above). Beyond the T-end detection issue, additional concerns exist:

T-peak search window is too short:

The T-peak search is 200-400 ms after R. At slow HR (50 BPM, RR = 1200 ms), the T wave peak can occur up to 550 ms after R. The window should be adaptive based on RR interval. A clinically sound rule is: T-peak search window = 0.2 * RR to 0.7 * RR, capped at 650 ms.

# Adaptive T-wave window based on RR interval

rr_ms = median_rr  # from HR calculation

t_search_start = r_idx + int(0.20 * rr_ms / 1000 * sampling_rate)

t_search_end   = r_idx + int(min(0.70 * rr_ms / 1000 * sampling_rate, 0.65 * sampling_rate))

U-wave contamination:

U waves (common in hypokalemia, bradycardia) appear 100-200 ms after T-peak. If U waves are present, the tangent method will correctly end at T-end. But a naive maximum search in the T window will sometimes pick U as T, inflating QT by 100-200 ms. Add a U-wave exclusion: if two peaks are detected in the T window, the first is T and the second is U — exclude the second from QT measurement.

QTc formula selection:

Bazett's formula (used in your code) over-corrects at HR > 100 BPM and under-corrects at HR < 60 BPM. For a real-time monitor where patient HR may vary widely, consider offering Fridericia as an alternative. The WHO currently recommends Fridericia for clinical studies. The formula change is trivial:

# Fridericia (better at extremes of HR):

qtc_sec = qt_sec / (rr_sec ** (1.0/3.0))  # Cube root instead of square root

# Framingham (best for ambulatory / Holter):

qtc_ms = qt_ms + 154 * (1 - rr_sec)

## 3.5 P Wave Duration

Once Bug #2 (peak_rel undefined) is fixed, the P duration algorithm structure is reasonable. However the 20% amplitude threshold for P wave acceptance is clinically too aggressive. In elderly patients, patients with atrial fibrillation history, and patients with electrolyte abnormalities, P wave amplitude can be as low as 0.05 mV (50 uV). With a 20% of R threshold, these low-amplitude P waves will always be rejected.

Recommended threshold adjustment:

# Original - too aggressive:

if p_amp > 0.2 * r_amp:

p_peaks.append(p_idx)

# Better - absolute minimum plus relative check:

p_amp_mv = (filtered_signal[p_idx] - baseline) / adc_scale  # Convert to mV

if p_amp_mv > 0.05 or p_amp > 0.10 * r_amp:  # 0.05 mV absolute OR 10% relative

p_peaks.append(p_idx)

# 4. Value Stability in Real-World Monitoring

## 4.1 Current Stability Status by Parameter

| Parameter | Currently Stable? | Smoothing Used? | What will cause instability in real patients |
| --- | --- | --- | --- |
| HR (BPM) | YES ✅ | EMA + Median | PVCs, AF, noise spikes — partially mitigated by median buffer |
| PR Interval | NO ⚠️ | None | Any variation in P wave detection — no beat averaging applied |
| QRS Duration | NO ⚠️ | None | Noise near Q/S detection, lead artifact — median of all beats but no quality gating |
| QT Interval | NO ⚠️ | None | T-end detection is highly noise-sensitive, no temporal smoothing |
| QTc | NO ⚠️ | None | Inherits QT instability; also varies beat-to-beat with RR variation |
| P Duration | NO ⚠️ | None | Low amplitude P waves missed/found alternately — high jitter expected |

## 4.2 Universal Stabilization Pattern to Apply to All Parameters

All interval parameters (PR, QRS, QT, P duration) should use the same stabilization pattern that was successfully applied to HR. The following template should be applied to each:

# Universal Stabilization Pattern (apply to PR, QRS, QT, P duration)

from collections import deque

_interval_buffers = {}   # Global or class-level

def stabilize_interval(value_ms, param_name, buffer_size=8,

physiological_min=0, physiological_max=9999):

"""

Apply median + EMA stabilization to any ECG interval parameter.

Same pattern as HR smoothing, tuned for slower-changing intervals.

"""

if param_name not in _interval_buffers:

_interval_buffers[param_name] = {'buf': deque(maxlen=buffer_size), 'ema': None}

state = _interval_buffers[param_name]

# Reject physiologically impossible values before adding to buffer

if not (physiological_min <= value_ms <= physiological_max):

# Return current EMA if we have one, else 0

return int(state['ema']) if state['ema'] is not None else 0

state['buf'].append(value_ms)

# Median of buffer (outlier-resistant)

median_val = np.median(list(state['buf']))

# EMA on top of median (slow alpha = stable display)

if state['ema'] is None:

state['ema'] = median_val

else:

alpha = 0.3 if abs(median_val - state['ema']) > 5 else 0.10

state['ema'] = (1 - alpha) * state['ema'] + alpha * median_val

return int(round(state['ema']))

# Usage:

stable_pr  = stabilize_interval(raw_pr_ms,  'pr',  buffer_size=8, physiological_min=80,  physiological_max=300)

stable_qrs = stabilize_interval(raw_qrs_ms, 'qrs', buffer_size=8, physiological_min=40,  physiological_max=200)

stable_qt  = stabilize_interval(raw_qt_ms,  'qt',  buffer_size=8, physiological_min=200, physiological_max=700)

stable_p   = stabilize_interval(raw_p_ms,   'p',   buffer_size=8, physiological_min=40,  physiological_max=200)

## 4.3 Signal Quality Index (SQI) Gating

This is the single most impactful change for real-world monitoring stability. Before updating any displayed parameter, compute an SQI and gate the update. If the signal is noisy, show the last valid value — do not update with garbage.

def compute_sqi(signal, sampling_rate=500):

"""

Simple Signal Quality Index: 0.0 (unusable) to 1.0 (clean).

Based on kurtosis of the filtered signal — ECG has high kurtosis

(spiky R peaks), noise is low kurtosis.

"""

from scipy.stats import kurtosis

if len(signal) < 200:

return 0.0

k = kurtosis(signal)

# ECG kurtosis typically 5-30, noise < 3, saturation > 50

if k < 3 or k > 50:

return 0.0

return min(1.0, (k - 3) / 20.0)

# Gate all metric updates:

sqi = compute_sqi(filtered_lead2)

if sqi > 0.3:  # Only update if signal quality is acceptable

hr = calculate_heart_rate_from_signal(lead2_data, 500)

pr = calculate_pr_interval(...)

# etc.

# else: keep last displayed values — do not update

# 5. Signal Filter Chain Audit

## 5.1 Current Filter Chain

The documented filter chain order is: ADC Normalization → Powerline Removal (50 Hz) → Baseline Correction → EMG Suppression → QRS Sharpening (disabled).

This order is correct and matches the IEC 60601-2-47 standard for ambulatory ECG processing. The key insight your team got right is applying the notch filter BEFORE baseline wander removal — many amateur implementations do this backwards, which allows 50 Hz harmonics to be spread into the baseline by the median filter.

## 5.2 Filter Parameter Verification

| Filter | Your Parameters | Industry Standard | Verdict |
| --- | --- | --- | --- |
| Bandpass (display) | 0.5–40 Hz, 4th order Butterworth | 0.05–40 Hz (AHA) or 0.5–40 Hz (monitor) | ✅ Correct for monitoring. For diagnostic 12-lead use 0.05 Hz low end. |
| Baseline wander | 120 ms median + 800 ms moving avg | 200 ms median + 600 ms mean (AHA) | ✅ Very close. Consider 200 ms median window for better QRS artifact rejection. |
| Notch filter | 50 Hz, Q=25, IIR notch | 50 Hz, Q=30–50 recommended | ⚠️ Q=25 gives wider notch band (~2 Hz). Increase to Q=35 to preserve nearby spectral content. |
| EMG low-pass | 35 Hz, 4th order Butterworth | 40 Hz (AHA diagnostic) | ✅ 35 Hz is acceptable for monitoring. 40 Hz for diagnostic quality. |
| QRS sharpening | Disabled by default | Not standard in ECG monitors | ✅ Correct decision to disable. Keep disabled. |

## 5.3 Important Note on filtfilt() in Real-Time Use

Your code consistently uses filtfilt() (zero-phase filtering) which is the correct choice for offline / batch processing because it eliminates phase distortion. However, filtfilt() is non-causal — it requires the entire signal buffer to compute. For real-time streaming, this means you are always filtering the full buffer window each update cycle.

This is acceptable if your buffer is a fixed sliding window (e.g., 2-5 seconds). However if your buffer grows unboundedly, filtfilt() will become slower over time. Ensure your signal buffers are fixed-length sliding windows, not appending lists.

# CORRECT: Fixed sliding window

BUFFER_SIZE = 2500  # 5 seconds at 500 Hz

signal_buffer = deque(maxlen=BUFFER_SIZE)  # Auto-drops oldest samples

signal_buffer.extend(new_samples)

filtered = display_filter(np.array(signal_buffer), fs)

# WRONG: Growing buffer

all_samples = []

all_samples.extend(new_samples)  # Grows forever, filtfilt slows down

# 6. Lead-Off Detection — Full Implementation Guide

## 6.1 Current Status

The detect_lead_off() function is written and correct in principle but is NOT integrated into the metric calculation pipeline. This means that if a lead electrode falls off, the system will continue calculating HR, PR, QRS, QT from the noise/flatline signal and display erroneous values. This is a patient safety concern for any clinical deployment.

## 6.2 Threshold Calibration Issue

The current high-variance threshold (variance > 10,000) is too permissive. At 500 Hz with 12-bit ADC, a lead fully off (showing random noise) would have a variance in the range of 50,000-500,000 ADC counts squared. The threshold of 10,000 corresponds to roughly ±100 ADC counts of fluctuation, which is actually within the range of a real ECG signal. This means high-noise but connected leads will NOT be flagged as lead-off.

Calibrated thresholds for 12-bit ADC at 500 Hz:

# Calibrated for 12-bit ADC (0-4096) at 500 Hz

LEAD_OFF_THRESHOLDS = {

'amplitude_min': 20,      # < 0.1 mV p-p → flatline/disconnected

'amplitude_max': 3000,    # > 14.6 mV p-p → saturation

'variance_min':  4,       # < variance → flatline (disconnected)

'variance_max':  250000,  # > variance → full noise (disconnected)

'saturation_lo': 10,      # ADC < 10 → rail low

'saturation_hi': 4085,    # ADC > 4085 → rail high

}

## 6.3 Integration into Pipeline (Critical Missing Step)

# In your main update loop (twelve_lead_test.py or equivalent):

lead_quality = {}

for lead_name, lead_data in all_lead_data.items():

is_off = detect_lead_off(lead_data, sampling_rate=500)

lead_quality[lead_name] = 'OFF' if is_off else 'OK'

# Gate metric calculation — only use Lead II if it's connected

if lead_quality.get('II', 'OFF') == 'OK':

hr = calculate_heart_rate_from_signal(lead_data_II, 500)

else:

hr = last_valid_hr  # Hold last known good value

display_alert('Lead II disconnected — HR unreliable')

# For each displayed lead plot:

for lead_name in all_leads:

if lead_quality[lead_name] == 'OFF':

show_lead_off_overlay(lead_name)  # Red 'LEAD OFF' text on plot

skip_drawing_waveform(lead_name)

# 7. Prioritized Fix Checklist

Apply fixes in this order. Items marked CRITICAL should be resolved before any clinical use or demonstration.

| # | Priority | Parameter | Fix Required | File |
| --- | --- | --- | --- | --- |
| 1 | CRITICAL | QT / QTc | Replace T-peak with T-end detection using tangent method | intervals.py |
| 2 | CRITICAL | P Duration | Fix undefined peak_rel variable — causes NameError / always returns 0 | expanded_lead_view.py |
| 3 | CRITICAL | Lead-Off | Integrate detect_lead_off() into metric calculation pipeline — gate all metrics | twelve_lead_test.py |
| 4 | WARNING | HR | Add ectopic beat (PVC) rejection — filter RR intervals >20% from median before HR calc | heart_rate.py |
| 5 | WARNING | PR Interval | Change PR detection from P-peak to P-onset; change floor from 120 to 80 ms (catches WPW) | intervals.py |
| 6 | WARNING | QRS Duration | Widen Q search window from 40 ms to 60 ms; change range floor from 60 to 40 ms | intervals.py |
| 7 | WARNING | QT | Make T-wave search window adaptive: use 0.2*RR to 0.7*RR instead of fixed 200-400 ms | intervals.py |
| 8 | WARNING | All Intervals | Apply median+EMA stabilization buffer to PR, QRS, QT, P duration (same as HR) | intervals.py |
| 9 | WARNING | All Metrics | Implement SQI gating — do not update display when signal quality is poor | heart_rate.py / new sqi.py |
| 10 | INFO | Lead-Off | Recalibrate variance threshold from 10,000 to 250,000 for 12-bit ADC | lead_off_detection.py |
| 11 | INFO | P Detection | Lower P amplitude acceptance threshold from 20% to 10% of R, add 0.05 mV absolute floor | intervals.py |
| 12 | INFO | Filters | Increase notch filter Q from 25 to 35 for sharper 50 Hz rejection | ecg_filters.py |
| 13 | INFO | QTc | Offer Fridericia formula alongside Bazett; use Fridericia when HR < 55 or > 100 BPM | intervals.py |
| 14 | INFO | Buffers | Verify all signal buffers use deque(maxlen=N) — not growing lists — to keep filtfilt() fast | All files |

# 8. What Your System Does Well

It is important to acknowledge the significant strengths of this codebase before summarizing the issues. Many of these design decisions reflect genuine engineering maturity:

- The multi-strategy R-peak detection approach (conservative/normal/tight) with automatic selection by standard deviation is an original and effective design. Most open-source ECG libraries use a single fixed strategy.

- The decision to disable variance normalization is exactly right. Many ECG systems fail in clinical use because of this mistake and your team identified and fixed it proactively.

- The Gold Standard baseline wander removal (120 ms median + 800 ms moving average) matches what commercial monitors like Philips IntelliVue and GE CARESCAPE use internally.

- Applying the notch filter before baseline wander removal is the correct order and many implementations get this wrong.

- The adaptive EMA alpha (0.5 for large changes, 0.10 for small) for HR smoothing is a well-tuned design that prevents both lag and jitter.

- The false asystole guard (minimum packet count before alarming) is an important clinical safety feature that is often missed.

- Disabling QRS sharpening by default is the correct conservative choice.

- Zero-phase filtering (filtfilt) throughout is correct for display quality.

# 9. Conclusion

The two most urgent issues are the T-end detection bug (which causes QTc to be systematically ~100 ms too short — a clinically dangerous error) and the peak_rel NameError in P duration calculation (which causes P duration to always return 0). These should be addressed immediately.

The lead-off detection integration is the next priority — without it, disconnected leads silently produce erroneous metrics with no indication to the user.

After those three are resolved, apply the universal stabilization buffer pattern to all interval parameters and implement SQI gating. These two changes will dramatically improve display stability in real patients with motion artifact, PVCs, and electrode noise.

The system architecture is sound and the existing fixes show good engineering judgment. With the corrections detailed in this report, this system can achieve clinical-grade reliability.

— End of Audit Report —

Generated by Claude (Anthropic) | February 12, 2026